## Slide 1: Antenna Course 23-Apr-2015 Lecture 6
Dr. Andrés Alayón Glazunov

## Slide 2: Lecture 6 Characterization of antennas in multipath
  MIMO and MIMO characterization
  Measurements in reverberation chamber
    Characterization of active terminals: TRP, TIS, Throughput data rate
  System Modelling
    Threshold Receiver: Probability of Detection (PoD), throughput

## Slide 3: Learning goals (after lab 1)
Describe how reverberation chambers work.
Describe how reverberation chambers can be used to measure MIMO antenna performance in rich isotropic multipath
Describe how reverberation chamber can be used to measure active terminals: Calibration of chamber and measurements

## Slide 4: Lecture 6 Characterization of antennas in multipath
  MIMO and MIMO characterization
  Measurements in reverberation chamber
    Characterization of active terminals: TRP, TIS, Throughput data rate
  System Modelling
    Threshold Receiver: Probability of Detection (PoD), throughput

## Slide 5: What is MIMO?
MIMO = Multiple Input Multiple Output
Multiple transmit (Tx) antennas & Multiple receive (Rx) antennas are used to produce multiple data streams
4G and Beyond (e.g., 5G) Communications Systems

## Slide 6: Shannon’s formula (1948) for the capacity of a SISO communication link in additive white noise channel (AWGN)
LOS case:
This formula gives the maximum achievable spectral efficiency in AWGN
Power and Bandwidth are resources available to improve communication

## Slide 7: How to use four antennas best in LOS according to Shannon?
Two separate (uncoupled) channels:

## Slide 8: How to use four antennas best in LOS?
Two separate (uncoupled) channels (2 bit streams) dividing power between them:
Use antennas in array for increased gain (1 bit stream):

## Slide 9: How to use four antennas best in LOS Use as many bit streams as “possible”, depending on available power.
How can we do this in practice?
Orthogonally polarized antennas
Spatial (or polarization) multiplexing.
2 bit streams (two 1x1 systems):
1 bit stream (classical array):

## Slide 10: MIMO system
Multiple Input Multiple Output
Transmitting and receiving on many antennas or antenna ports
Many radio sub channels
Max available capacity increases when adding antenna elements
We must measure (estimate) the channel before using it for data transfer at the receiver
Can send many bit streams by using spatial multiplexing (diagonalization of channel matrix). Channel must be known at the transmitter too.
nt = 1, 2, …M antenna ports
nr = 1, 2, …N antenna ports
Rich scattering environment causing Rayleigh fading
Tx ports
Rx ports
H is the channel matrix containing all “S21”-parameters between all Tx and all Rx ports

## Slide 11: MIMO system in multipath environment, Shannon’s extended formula (CSI = Channel state information)
No CSI at transmitter, full CSI at receiver, power equally divided:
H is the channel matrix containing all S-parameters between all Tx and all Rx ports
nt = 1, 2, …M antenna ports
nr = 1, 2, …N antenna ports
Rich scattering environment causing Rayleigh fading
Tx ports
Rx ports

## Slide 12: Result measuring the channel measurements and using channel capacity formula for 3x6 MIMO. Effect of mutual coupling/correlation.
3x6 MIMO
Reduction in capacity for small element spacing d is due to reduction in embedded element efficiency of each monopole, and correlation between channels
d
d =

## Slide 13: MIMO capacity for SNR = 15 dB antenna modeled in IE3D and and external RIMP ray model. Measurements in reverberation chamber.

## Slide 14: MIMO concepts in modern digital communication systems
MIMO (and communications in general) requires an estimate (measurement) of the channel matrix H.
For small SNR in rich isotropic multipath:
  Antenna diversity by Maximum Ratio Combination (i.e. superposition by complex conjugate channel weighting)
For large SNR
  MIMO multiplexing/multi-stream may produce two or more bit streams at same frequency to drastically multiply channel capacity
  Processing involves diagonalization of the channel matrix (eigenvector formation). Each eigenvector represents a “bit stream channel” (channel must be known at both the receiver and the transmitter).
For large SNR in LOS environments:
  MIMO reduces to classical beam-forming, 2 beams (bit streams) can be formed when using 2x2 MIMO if the ports excite orthogonal polarizations.
  We can also achieve 2 bit streams by using co-polar elements with large separation (not practical).
For large SNR in advanced systems
  MultiUser MIMO, i.e. many send different bit streams at same frequency to different users in same environment, Massive MIMO.

## Slide 15: Lecture 6 Characterization of antennas in multipath
  MIMO and MIMO characterization
  Measurements in reverberation chamber
    Characterization of active terminals: TRP, TIS, Throughput data rate
  System Modelling
    Threshold Receiver: Probability of Detection (PoD), throughput

## Slide 16
Traditional characterization method: Anechoic Chambers

## Slide 17: Reverberation Chamber

## Slide 18: Simulated Electrical Field in cavity during stirring
In-house MoM code
Cavity Green’s functions

## Slide 19: Excitation of multiple resonances in a reverberation chamber……... cavity modes

## Slide 20: Excitation of multiple resonances due to stirring (moves the resonances) and the finite bandwidth of each mode
Df
0
6

## Slide 21: TE and TM-modes in rectangular cavity
where na=0, 1, 2, …, nb=0, 1, 2, …
and nc=1, 2, 3, …
Each mode corresponds to 8 plane waves:

## Slide 22: Angles of incidence of plane waves in 0.8m x 1.1m x 1m chamber (GSM 900 band)
From Kent Rosengren, Intenna AB
Conclusion: Reverberation chambers corresponds to isotropic multipath environment with uniform distribution of wave directions in both azimuth and elevation.

## Slide 23: Field statistics similar to rich isotropic multipath propagation environment
In reverberation chamber: 1 cavity mode = 8 plane waves.
Angles of incidence uniformly distributed
In real environment: Some elevation weighting

## Slide 24: Excitation of multiple resonances due to stirring (moves the resonances) and the finite bandwidth of each mode
Df
0
6

## Slide 25: Mode bandwidth (D. Hill)
The mode bandwidth can be divided into four contributions
  wall losses
  leakage
  losses due to antennas
  lossy objects
Mode bandwidths are additive

## Slide 26: Example of mode bandwidths
Estimated mode bandwidth for reverberation chamber (1x0.8x1.6m
3
)
850
900
950
1000
-1
0
1
2
3
4
5
Frequency (MHz)
Bandwidth (MHz)
1600
1700
1800
1900
2000
-1
0
1
2
3
4
5
Frequency (MHz)
Bandwidth (MHz)
Total
Lossy objects
Leakage
Antenna load
Wall losses

## Slide 27: Mode bandwidth and Hill’s Transfer function
The mode bandwidth is related to the transfer function of a chamber by (from Hill, 1994)
Thereby we can measure radiation efficiency and radiated power
Compare with Friis’ for LOS:
(
)

## Slide 28: The field inside the chamber
The test antenna will see a field inside the chamber that is uniform and isotropic in the statistical sense
Standard deviation for the mean received power goes as (result of complex Gaussian process):
Nind is the number of statistically independent samples.
This is a general formula for Gaussian processes.

## Slide 29: Conf. interval for mean received power
0
50
100
150
200
-10
-8
-6
-4
-2
0
2
4
Number of independent samples
Conf. interval (dB)
95% (2
s
)
99%
The number of independent samples is determined by how well we stir the modes in the chamber, and how many modes that are excited (i.e. chamber size and frequency).

## Slide 30: Number of independent samples Nd
The number of independent samples is determined primarily by the mode density in the chamber, i.e., by the number of modes per MHz.
It depends on chamber loading, mode stirring methods, mechanical stirrer shapes, and chamber shape (at least for small chambers).

## Slide 31: Reverberation chamber and set-up for antenna measurements
Stirring methods: mechanical, platform, polarization and frequency.
S22
S11
The channel is S21

## Slide 32: Cumulative probability distribution of fading
S21 is complex Gaussian distribution and envelope Rayleigh
S21
Abs(S21)

## Slide 33: Coherence bandwidth and the inverse time delay spread can be controlled by loading Coherence bandwidth = mode bandwidth

## Slide 34: RTS chamber and laptop with USB dongle for LTE with OFDM

## Slide 35: Data extraction for reflection coefficient and radiation efficiency (averaging over stirring positions and frequency)
Radiation efficiency (ref antenna is assumed to have 100% eff):
Pref is relative received power of reference antenna
Free space reflection coefficient of AUT:
Relative received power of AUT:

## Slide 36: Measurements of free space antenna impedance Example: Dipole 10 mm from cylinder surface

## Slide 37: Data extraction for reflection coefficient and radiation efficiency (averaging over stirring positions and frequency)
Radiation efficiency (ref antenna is assumed to have 100% eff):
Pref is relative received power of reference antenna
Free space reflection coefficient of AUT:
Relative received power of AUT:

## Slide 38: Validation case: Measured and computed radiation efficiency of dipole close to lossy cylinder
Dotted lines: simulation by FDTD
Inaccurate due to size of chamber (present chamber is larger and accurate to lower frequency)
Notes: Lägg in figur med mätning av strålningseffektivitet, innehållande erad, Pref och Paut från föregående OH-bild.

## Slide 39: STD Uncertainty of efficiency in Bluetest chambers (RIMP)
With head phantom
Without head phantom

## Slide 40: Accuracy
Bluetest states STD = 0.5 dB
Results of several benchmarking and Round Robin tests:
  Measurements of efficiency and radiated power in reverberation chamber show better repeatability than measurements in many anechoic chambers
Extensive benchmarking has been done by Antenna Center of Excellence (ACE, a European Network of Excellence)

## Slide 41: Diversity measurements

## Slide 42: Cumulative probability distributions and effective and apparent diversity gain

## Slide 43: Setup for measuring MIMO antennas (up to 3x6 MIMO channels)
Fixed Wall Antenna
Reference Dipole
6 Element Circular
MIMO Array

## Slide 44: Lecture 6 Characterization of antennas in multipath
  MIMO and MIMO characterization
  Measurements in reverberation chamber
    Characterization of active terminals: TRP, TIS, Throughput data rate
  System Modelling
    Threshold Receiver: Probability of Detection (PoD), throughput

## Slide 45: Measurement setup – TRP
Phone output power and traffic channel controlled by base station simulator
Power is measured for each
  frequency
  excitation antenna
  stirrer position

## Slide 46: Measured radiated power of 20 phones in 5 positions (free space + CENELEC)

## Slide 47: Example of results for TRP (TCP) 20 phones, 5 positions, 900 MHz

## Slide 48: Lecture 6 Characterization of antennas in multipath
  MIMO and MIMO characterization
  Measurements in reverberation chamber
    Characterization of active terminals: TRP, TIS, Throughput data rate
  System Modelling
    Threshold Receiver: Probability of Detection (PoD), throughput

## Slide 49: Comparison between ideal threshold receiver and practical LTE device
SISO case. Conductive measurements with cable.
New theory for throughput is based on using an ideal threshold receiver.

## Slide 50: Threshold Receiver Model
The throughput in conducted measurements can be modelled as

## Slide 51: Threshold Receiver Model in Fading Channel RIMP
Throughput is averaged over fading states
Model
Measurements

## Slide 52: Theory and measurements of LTE device: Includes MIMO and OFDM diversity
Solid curve: Measured in Bluetest RTS chamber
Dotted curves: Theory with threshold receiver model
1x1 SISO
1x2 SIMO
& with
OFDM
threshold
Average power [dBm]
Relative throughput
1x1 SISO
1x2 SIMO
& with
OFDM
threshold
Average power [dBm]
Relative throughput

## Slide 53: Relation between Tput and CDF in RIMP
Diversity gain is clearly seen in both Tput and CDF curves,
representing cumulative improvement of 1% worst users

## Slide 54: Details about theoretical model in RIMP Note the inverse relation to Pav comparing Tput and CDF Normally we determine and plot CDF(Pav)
Relationships for powers in dBm:

## Slide 55: Popularity of “OTA” tests in RIMP OTA = Over-The-Air

## Slide 56: Our present research: includes also random-LOS & testing the real-life hypothesis
Environment | Equivalent measurement method | Antenna quality measure | MIMO and diversity capability
Free space (pure-LOS) | Anechoic chamber | Deterministic case: Realized gain User-random case: Not known | To some degree
Real-life environments | Expensive drive tests | No unique quality measure | Yes
Rich isotropic multipath (RIMP) | Reverberation chamber | Total radiation efficiency erad & Diversity gain in dBR & Multiplexing efficiency in dBiid | Yes
Hypothesis: If wireless terminals work well in RIMP and random pure-LOS environments, they will work well also in real-life environments. It remains to be proven.

## Slide 57: Learning goals (after lab 1)
Describe how reverberation chambers work.
Describe how reverberation chambers can be used to measure MIMO antenna performance in rich isotropic multipath
Describe how reverberation chamber can be used to measure active terminals: Calibration of chamber and measurements